pH METERS

PURDUE UNIVERSITY INSTRUMENT VAN PROJECT

BECKMAN pH METER INSTRUCTIONS AND REMINDERS

(1-8-96)

Buffers (pH = 4, 7, and 10) MUST be available at the school site.  Virtually all buffers must be diluted.  READALL LABEL INSTRUCTIONS!

These instructions are for the use of BECKMAN FUTURA PLUS ELECTRODES only!

You have been supplied with a storage case with test tubes in which to store the electrodes between uses.  DO NOT remove the test tubes from the storage case.  The test tubes are filled with storage solution (4M KCl saturated with AgCl), which has been provided.  DO NOT add any liquid to these test tubes.

Remove the electrode from the test tube.  Remove the plug from the fill hole.  The fill hole cover stops or reduces the flow of filling solution from the elctrode during shipment or during periods of non-use.  It must be removed when the electrode is in use to permit free flow of filling solution through the reference junction.  Fill the electrode with filling solution.  Be sure that you use the filling solution provided.

Rinse the bulb and junction area clean with deionized water.  Connect the keeper cable to the electrode and pH meter.   In order for the meter to operate properly the BNC connector must fit snugly.  To connect the probe to the meter:

1.  Remove the protective cover from the input post.

2.  Place the BNC connector over the input post and push down.

3.  Turn clockwise to complete the connection.

Electrodes should never be left "high and dry".  They should always be immersed except when they are being transferred from one solution to another.  You have been supplied with a storage case in which to store the electrodes between uses.  Never store the electrode in distilled or deionized water.  This may lead to slow sluggish response caused by junction clogging.  Between measurements immerse the electrode in a buffer solution (4 or 7 recommended).  When the electrode is not in use, it is highly recommended that it is stored with the bulb and reference junction immersed in the KCl/AgCl storing solution.

ALWAYS rinse the electrodes with distilled water and blot with a Kimwipe® upon removal from one solution before placing in another solution.

Glass electrodes are made up of very thin glass and are VERY FRAGILE.  Care should be taken not to bump them on the bottom or sides of a container or with a stir bar.

pH METERS

PURDUE UNIVERSITY INSTRUMENT VAN PROJECT

For use of the Beckman pH Meter in the Acid Range

You will need to calibrate two points: pH = 7 and pH = 4

To standardize (acid range), follow these steps:

1.   Turn ON (O) and CLEAR (C) instrument.

2.   Immerse electrode in standard 1, pH = 7 buffer, stir briefly.

3.   Press pH, then press STD.

4.   After the “eye” symbol stops flashing, the display will read pH of standard 1.

5.   Remove the electrode.  Rinse it with distilled water, and blot GENTLY with a 		      Kimwipe®.

6.   Immerse electrode in standard 2, pH = 4 buffer, stir briefly.

7.   Press STD.

8.   After the “eye” symbol stops flashing, the display will read pH of standard 2.

9.   Remove the electrode.  Rinse it with distilled water, and blot GENTLY with a 		      Kimwipe®.

10. Immerse electrode in sample solution, stir briefly.  Press pH.  After the “eye”

symbol stops flashing, the display will read pH of sample solution.

For Use of the Beckman pH Meter in the Base Range

You will need to calibrate two points: pH = 7 and pH = 10.

To standardize (base range), follow these steps:

1.   Turn ON (O) and CLEAR (C) instrument.

2.   Immerse electrode in standard 1, pH = 7 buffer, stir briefly.

3.   Press pH, then press STD.

4.   After the “eye” symbol stops flashing, the display will read pH of standard 1.

5.   Remove the electrode.  Rinse it with distilled water, and blot GENTLY with a 		      Kimwipe®.

6.   Immerse electrode in standard 2, pH = 10 buffer, stir briefly.

7.   Press STD.

8.   After the “eye” symbol stops flashing, the display will read pH of standard 2.

9.   Remove the electrode.  Rinse it with distilled water, and blot GENTLY with a 		      Kimwipe®.

10. Immerse electrode in sample solution, stir briefly.  Press pH.  After the “eye”                  symbol stops flashing, the display will read pH of sample solution.